AeroSense

Quick Guide

Portable Air Quality Monitor

# What It Does

AeroSense measures the air around it and shows how clean it is right now. The readings appear on the device immediately, and it works perfectly well on its own. Units that have been given Wi-Fi details also upload their readings to a private cloud database — see Wi-Fi and Cloud Sync.

| Reading | What it is |
| --- | --- |
| PM1.0 / PM2.5 / PM10 | Airborne particles — smoke, soot, dust, haze (µg/m³). PM2.5 matters most for health. |
| Gas (MQ-2) | Combustible gas and smoke — LPG, cooking gas, methane. A relative level, not a calibrated figure. |
| AQI | Air Quality Index, 0–500. Calculated from PM2.5 and PM10. The headline number. |

# Getting Started

- Place the device on a flat, dry surface at desk height, away from draughts and heat.

- Plug in the USB power. A start-up screen shows for about 1.5 seconds.

- The Home screen appears with a large AQI number.

- Wait 60 seconds while the gas sensor warms up. A high Gas reading during this time is normal — see below.

- Press Button 2 for the next page, Button 1 for the previous one. Hold Button 1 to jump back to Home.

# The First Minute

The MQ-2 contains a small heater. Until it reaches working temperature it reads close to maximum. This is normal hardware behaviour, not gas in the room.

AeroSense keeps sampling throughout, but switches gas warnings off for 60 seconds so you get no false alarms. AQI is unaffected and works normally, because it comes only from the particulate sensor.

While warming up, the Home screen shows warming in place of the gas number, the Gas Monitoring page shows WARMING UP and a countdown, and System Status shows WARMUP 45s. When done it shows READY.

Note: Warm-up runs once from power-on. Changing pages or pressing buttons does not restart it.

# The Two Buttons

A press is a quick tap. A hold is about half a second.

|  | Button 1 | Button 2 |
| --- | --- | --- |
| Press | Previous page | Next page |
| Hold | Go to Home | Open Quick Menu |

## In the Quick Menu

| Action | Result |
| --- | --- |
| Press Button 1 / Button 2 | Move the highlight up / down |
| Hold Button 2 | Confirm the highlighted option |
| Hold Button 1 | Cancel and close the menu |

Options: Reset Statistics, Wi-Fi Setup, Go to Home, Close Menu.

# The Seven Pages

Button 2 steps through them in this order and loops back to Home.

| # | Page | What you see |
| --- | --- | --- |
| 1 | Home | Large AQI number and category, plus PM1.0, PM2.5 and PM10 cards and a gas readout along the bottom. The bar under Gas fills up as the level rises. |
| 2 | Air Quality | The three PM readings with bars, and a trend symbol: + rising (red), – falling (green), = steady. |
| 3 | Gas Monitoring | MQ-2 with a colour gauge and a status word: LOW, MEDIUM, HIGH or DANGER. Below it, the highest reading seen this session. |
| 4 | Sensor Dashboard | All four sensor values in a 2×2 grid. The pulsing teal dot means the device is actively reading. |
| 5 | System Status | Sensor health, gas warm-up state, free memory and uptime. |
| 6 | History Graphs | The last two minutes of PM1.0, PM2.5, PM10 and Gas as line graphs. |
| 7 | Statistics | Min, average and max for PM2.5, PM10 and MQ2. Clear them with Reset Statistics. |

Note: Statistics start recording at power-on, so the MQ2 Max figure includes the high warm-up readings. Use Reset Statistics after warm-up for a meaningful gas number.

# Understanding the AQI

AQI turns the particle readings into one number from 0 to 500 — higher is always worse. AeroSense uses the standard US EPA scale, the same one used by weather apps, so the number is directly comparable.

It calculates an index from PM2.5 and another from PM10, then shows the worse of the two. The gas sensor is not included, because its reading is relative rather than a measured concentration.

| AQI | Shows | What it means and what to do |
| --- | --- | --- |
| 0–50 | GOOD | Air is fine. No action. |
| 51–100 | MODERATE | Acceptable. Very sensitive people may notice. |
| 101–150 | SENSITIVE | Unhealthy for children, older adults, asthma and heart conditions. Ventilate or filter. |
| 151–200 | UNHEALTHY | Everyone may feel effects. Reduce activity, ventilate or run a purifier. |
| 201–300 | V.UNHEALTHY | Health warnings for everyone. Avoid exertion, act on the source. |
| 301–500 | HAZARDOUS | Emergency conditions. Leave the area if you can. |
| — | - - - / NO DATA | The particle sensor is not reporting yet. Normal for the first few seconds. |

# Warnings

A banner interrupts the screen for about 3 seconds, then clears itself. Nothing to dismiss. If a warning keeps coming back, the condition is still there.

| Banner | What to do |
| --- | --- |
| Smoke Detected! | Ventilate. Check for a gas leak, unattended cooking or burning. Leave the room if the cause is not obvious. |
| Poor Air Quality | Ventilate or filter the air. Look for a source such as cooking, smoke or dust. |

Gas warnings are off during the first 60 seconds after power-on.

# Wi-Fi and Cloud Sync

This is optional. A unit with no Wi-Fi details stored never switches its radio on and sends nothing — it is a standalone monitor, and it shows no connection indicator.

When Wi-Fi is set up, AeroSense uploads its current readings every 5 seconds to a private Google Firebase database, and adds a history record once a minute. It sends the particulate and gas readings, the AQI and its category, its uptime and signal strength, and a timestamp — nothing else. Only the single account the device signs in with can read or write that data.

The small four-bar meter at the top right shows how it is getting on. The bars are the Wi-Fi signal strength. The colour is the sync state: green means readings are reaching the database, amber means it is still connecting or signing in, red means uploads are failing, and grey means there is no network.

To put AeroSense on a network, hold Button 2 to open the Quick Menu and choose Wi-Fi Setup. The monitor raises its own network, AeroSense-Setup, with the password shown on the screen. Join it from a phone and a page opens listing the networks in range: tap one, type its password, and the monitor connects and remembers it. Sync trouble never affects the readings on screen.

# If Something Looks Wrong

| Symptom | What to do |
| --- | --- |
| Blank screen | Check the USB cable and power supply. Charge-only cables are a common culprit. |
| Gas reading very high just after switch-on | Normal warm-up. Wait for the countdown. |
| Gas still high long after warm-up | Ventilate and watch it. If it stays high, leave the device powered on for a few hours — these sensors settle further with use. |
| AQI stuck on - - - | Check System Status. If PMS5003 shows NO SIGNAL, check the sensor connection and that its fan is running, then power off and on. |
| Screen frozen, teal dot not pulsing | Unplug power, wait five seconds, plug back in. |
| Buttons seem unresponsive | A warning banner holds the page for about 3 seconds. Wait for it to clear. |

# Safety and Care

- Indoors only. Keep dry and away from steam, direct sunlight and heat.

- Do not cover the device or block its air intake.

- Clean with a dry cloth only. Do not open the enclosure.

- For steadier gas readings, leave it powered on continuously.

Please read the following before relying on this device.

- AeroSense is an indicative instrument, not a certified safety device. It is not a substitute for an approved smoke alarm, CO alarm or gas detector, and must not be used as one. Fit approved alarms as your local regulations require.

- The gas sensor is uncalibrated. It shows change, not measured concentrations, and responds to more than one gas.

- It is not a medical device. Do not use its readings to make medical decisions.

- If you suspect a gas leak or carbon monoxide, act on that suspicion, not on this device. Leave the area, ventilate, and call your gas supplier or emergency services. Carbon monoxide is colourless, odourless and can be fatal — and this device does not detect it.

AeroSense — Quick Guide · Keep for reference